## Slide 1: Image Pipeline for DHM
Fabian Braun
Master Thesis
07. September 2023, Zurich

## Slide 2: Research Goal: how do bacteria accumulate biomass
Goal of any organism: maximize biomass based on available resources.
Models for bacteria are based on biomass, but often only length is measured.
L
w
Length
Volume
Mass
Questionable assumption that density is constant:
Oldewurtel, 2021
Literature assumption:
Notes: Label-free imaging: No contrast agant or dye needed -> living cells
Non-invasive: Does not damage or alter imaged samples -> suitable for long-term observations of living cells and tissues
High Sensitivity to Thickness -> sutable for studying cell morphology
Quantitative measurements: quatitative information about optical path length changes -> cell volume, dry mass, intracellular refractive index
High-throughput imaging: Can be adepted for high-throuput imaging systems, allowing for analysis of large sample populations and fast changing samples
Three-dimensional reconstruction: True 3 dimensional imaging by taking several images from all sides, pseudo-3d, by changing the focus distacne -> tracking
Multimodal imaging: Can be combined with other imaging modalities such as fluorescence microsopy or brightfield imaging- > comprehensive view of the sample‘s properties

## Slide 3: Measuring dry mass with the phase difference image
Mass can be measured by measuring the phase difference.
Generally the denser the material the slower the speed of light
Coherent light source
Sample
Cell
2π
Optical path difference (OPD)
Phase difference =
= wavelength
Methods measuring the phase shift are called quantitative phase imaging (QPI).
Notes: Label-free imaging: No contrast agant or dye needed -> living cells
Non-invasive: Does not damage or alter imaged samples -> suitable for long-term observations of living cells and tissues
High Sensitivity to Thickness -> sutable for studying cell morphology
Quantitative measurements: quatitative information about optical path length changes -> cell volume, dry mass, intracellular refractive index
High-throughput imaging: Can be adepted for high-throuput imaging systems, allowing for analysis of large sample populations and fast changing samples
Three-dimensional reconstruction: True 3 dimensional imaging by taking several images from all sides, pseudo-3d, by changing the focus distacne -> tracking
Multimodal imaging: Can be combined with other imaging modalities such as fluorescence microsopy or brightfield imaging- > comprehensive view of the sample‘s properties

## Slide 4: Why are we interested in quantitative phase imaging (QPI)?
Digital holographic Microscopy (DHM) is one of many QPI methods.
Cells and bacteria are generally phase object, hardly visible by eye.
Advantages:
  Label-free imaging, non-invasive
  High Sensitivity to thickness
  Quantitative measurements
  High-throughput imaging
  Three-dimensional reconstruction
  Multimodal imaging
Intensity
Phase
Notes: Label-free imaging: No contrast agant or dye needed -> living cells
Non-invasive: Does not damage or alter imaged samples -> suitable for long-term observations of living cells and tissues
High Sensitivity to Thickness -> sutable for studying cell morphology
Quantitative measurements: quatitative information about optical path length changes -> cell volume, dry mass, intracellular refractive index
High-throughput imaging: Can be adepted for high-throuput imaging systems, allowing for analysis of large sample populations and fast changing samples
Three-dimensional reconstruction: True 3 dimensional imaging by taking several images from all sides, pseudo-3d, by changing the focus distacne -> tracking
Multimodal imaging: Can be combined with other imaging modalities such as fluorescence microsopy or brightfield imaging- > comprehensive view of the sample‘s properties

## Slide 5: Digital holograpic Microsopy (DHM)
BS
Phase image
Hologram
Source
L
BS
CCD
Scope
Sample
M
M
O
R
L
DHM acquisition
Hologram
Applications:
  Phase image
  3D tracking
R
O
Z
X
θ

## Slide 6: The pipeline consists of two parts
Phase image construction
Segmentation and lineage tracking
Hologram
Phase image
Construction
Segmentation and tracking

## Slide 7: DHM-Image reconstruction with LyncéeTec Koala
Hologram
Set focus
Intensity image
Phase image
Notes: Isolation of objective beam,
Propagation back to reconstruction distance

## Slide 8: Noise reduction methods
Phase images are noisy:
  Background noise, due to dust on optical setup
  Temporal noise, due to vibrations
  Optical noise, from speckles
  Spatial noise from debris in the sample
In multiple slighly shifted images:
  Background noise is static: isolated by taking the mean.
  Spatial averaging: aligning images and taking the mean. Reduces random noise (temporal and optical noise) by .
shifts
Cell culture
Close up
Images

## Slide 9: Changes to previous workflow
Hologram
Phase Image
Intensity Image
Avg. Phase Image
~2 MByte each
~1 MByte
Hologram
Avg. Phase Image
~1 MByte
~125 MByte per avg. Image
~25 MByte per avg. Image
Before
Now
Focusing
Set focus
Spatial avg.

## Slide 10: Process 1: surface plane subtraction
Uneven surface: Caused by imperfect alignment of optical instruments in DHM.
Subtraction with Koala only possible during experiment and not in postprocessing
Surface: Adjusted using linear regression on polynomial features of pixels distance.
Before
After
Difference
Plane subtr.
Focusing
Background subtr.
Spatial averaging

## Slide 11: Process 2: focusing
Deviation from set focus distance: Unknown reason and also unpredictable
Automatic focusing: Based on image features
Image depicted has a strong focus deviation.
Before
After
Difference
Plane subtr.
Focusing
Background subtr.
Spatial averaging

## Slide 12: Images with various reconstruction distances
Klar mache, das reconstruction und focus
Find region of focus with maxima of phase sharpness.
Find exact single global focus plane with minimum of std(amplitude).

## Slide 13: Process 3: background subtraction
Static Noise: Dust or other disturbances on the optical system
Median of unshifted image stack
Calculated on the first ten timesteps as later images can get croweded
Before
After
Difference
Plane subtr.
Focusing
Background subtr.
Spatial averaging

## Slide 14: Process 4: spatial averaging
Reduces random noise.
Calculates relative shifts between images using phase cross correlation.
Stacks shifted images, calculates the average, and then trims the outer portion.
Before
After
Difference
Plane subtr.
Focusing
Background subtr.
Spatial averaging

## Slide 15: Results: per pixel error
Data based on 1600 images, mean signal of a bacteria is ~25 nm.
Error measured on background, except for focusing  unfocused vs. focused
Few images are significantly out of focus.

## Slide 16: Results: focusing
Focus error has a none zero mean  considerable decrease in mass
Figure shows mass is decreased with n of 25 images 1cm out of focus.

## Slide 17: Pipeline part two
Phase image construction
Segmentation and lineage tracking
Hologram
Phase image
Construction
Segmentation and tracking

## Slide 18: Previous workflow for segmentation and tracking
Problems:
  Ilastik does not generalize well: each experiment has to be segmented.
  Dependent on human segmentation  not objective, especially over time
  A lot of manual work
Phase images
Segmentation
Interpretation
Tracking
Results
Ilastik

## Slide 19: Segmentation and Tracking with U-Net
U-Net is a neural network and the current standard in biological image processing.
Encoder-decoder structure.
Skip connections: preservation of spatial information and multilevel feature interaction
Data augmentation allows the model to generalize well even with few data.
Deep Learning for Time-Lapse Analysis (DeLTA) is based on the U-Net and used in this work.
Skip connections

## Slide 20: Preparation of training data: Segmentation with Ilastik
Train Ilastik by categorizing pixels into core, bacteria, boundary and background.
Prediction of other pixels based on preselected feature set
Separation between bacteria based on core.
Full size based on the bacteria
Manual pixel categorization
Ilastik‘s learned segmentation

## Slide 21: Training data augmentation methods
Original
Horzontal flip
Vertical flip
Rotation
Zoom
Gaussian blur
Gaussian noise
Elastic deformation

## Slide 22: Results: Time-lapse video of bacterial growth and linage
15 µm

## Slide 23: Output linage with features, cell cycles
2 µm

## Slide 24: Summary
An automatic focusing has been implemented
Storage space is reduced by a factor of 5
Automated segmentation and tracking with the help of a U-Net
Visual and easy to use output files
Easy to expand with other image types (brightfield, fluorescent)
Pipeline

## Slide 25: Phase image  Mass
Calculate optical path volume (minus bacteria volume * difference in refractive index) * 1/refractive increment
refractive increment from table, by knowing the composition of the bacteria

## Slide 26: How does a holgram work?
In 1948 D. Gabor found principle that captures both amplitutde and phase information of an object using interference pattern.
Gabor's hologram is essentially the record of this interference pattern between an object beam and a reference beam on a photosensitive medium.
When the developed hologram is lit with the same light, the object beam reappears.
This object beam can be followed back; focus is achieved when the object distance matches the difference between object and hologram planes.
britannica.com/technology/holography

## Slide 27: Outlook
  Improve training data generation, by consentrating on a smaller but more precise image batch.
  Multimodal approach: combining with fluorescent or brightfield images

## Slide 28: Pros and cons
Pros:
  Decrease in storage space by a factor of 5
  Automated focus distance detection
  No segmentation training needed for each induvidual experiment
  Easily to use and visual output files
  Pipeline combines all processes into one easy to use script
Cons:
  Post processing time is increased
  Segmentation accuracy is not improved